# LKAT MUT

# LKAT System Functional Test Set

# Description

The LKAT MUT ( Meter Unit Tester ) is a portable device that can be used together with a standard digital multi-meter to verify proper operation, scaling and alarm trip points of LKAT systems already installed in the field.

# Application

The LKAT MUT was developed as an accessory for LKAT systems to support ISO9000 and other quality system requirements which require periodic verification or testing of measurement / protection systems in the plant.  The LKAT MUT also enables end users to change alarm trip points if needed in the field

# Key Features

- Small, portable and powered by the LKAT under test.

- Allows users to test, verify and reset multiple aspects of DynAmp LKAT systems

- Supports methodology to verify proper measuring head operation and head to metering unit cable integrity

- All signal and alarm outputs are tested at actual metering unit output connectors to validate internal harnessing and connections.

# Measurement and Functional Tests

Measurement Head Performance / Measurement

Tests:	Internal Head electronics

Proper magnetic field measurement

Head to Meter Unit cabling

Internal harnesses and output connectors

Accuracy Diagnostics Function

Tests:	Detection functionality

Proper PCB LED operation

Proper Relay function

Internal harnesses and output connectors

Primary Alarm Relay

Tests:	Primary Alarm Trip setting

Proper PDB LED operation

Proper Relay function

Internal harnesses and connectors

Optional PE Alarm Relays ( protection extensions )

Tests:	Both PE Alarm Trip settings

Proper PCB LED operation

Proper Relay function

Internal harnesses and connectors

# Tools Needed

LKAT MUT System

4 ½ digit DMM with 0.05% basic DC accuracy

2 banana to banana plug test leads to connect MUT & DMM

LKAT Users Manual & MUT test form / Excel spreadsheet

LKAT serial no. specific calibration data from DynAmp

# Testing Methodology Overview

- Switch rectifier control to manual or alternative control

- With LKAT system operating and bus current between 20% and 100% of full scale, open metering unit door and measure head signals at PCB connector.  Also record LKAT measurement outputs

- Power down LKAT system, disconnect head and signal output cables from the meter unit, then connect the 3 cables from the LKAT MUT.  Power up LKAT system.

- Using the adjustments on the LKAT MUT, set the corresponding ‘head’ signal for the desired current level ( typically just before the alarm trip point )

- Using MUT adjustments, increase or decrease signal to check alarm trip points via alarm LED on LKAT PCB as well as the relay trip LED on LKAT MUT.

( Similar method used to modify alarm trip points if needed )