# LKAT MUT

# LKAT System Functional Test Set

# Description

The LKAT MUT ( Meter Unit Tester ) is a portable device that can be used together with a standard digital multi-meter to verify proper system operation, scaling, alarm trip points and adjustments of LKAT systems already installed in the field.

# Application

The LKAT MUT was developed to support ISO9000 requirements including periodic verification or testing of measurement / protection systems in the plant.  It is also an excellent tool to support system commissioning work in the field.

# Key Features

- Small, portable and powered by the LKAT under test.

- Allows users to test, verify and reset multiple aspects of DynAmp LKAT systems

- Supports methodology to verify proper measuring head operation and head to metering unit cable integrity

- All signal and alarm outputs are tested at actual metering unit output connectors to validate internal functions, harnessing and connections.

# Measurement and Functional Tests

Measurement Head Performance / Measurement

Tests:	Internal Head electronics

Proper magnetic field measurement

Head to Meter Unit cabling

Internal harnesses and output connectors

Accuracy Diagnostics Function

Tests:	Detection functionality

Proper PCB LED operation

Proper Relay function

Internal harnesses and output connectors

Primary Alarm Relay

Tests:	Primary Alarm Trip setting

Proper PDB LED operation

Proper Relay function

Internal harnesses and connectors

Optional PE Alarm Relays ( protection extensions )

Tests:	Both PE Alarm Trip settings

Proper PCB LED operation

Proper Relay function

Internal harnesses and connectors

# Tools Needed

LKAT MUT with included banana to Banana test leads

(adapters for LKAT OEM and SIGMA versions included)

LKAT Users Manual & MUT test form / Excel spreadsheet

LKAT serial no. specific calibration data from DynAmp

4 ½ digit DMM 0.05% basic DC accuracy (not included)

# Test Methodology Overview

- Switch rectifier control to manual or alternative control

- With LKAT system operating and bus current between 20% and 100% of full scale, open metering unit door and measure head signals on PCB connector.  Also record LKAT measurement outputs

- Power down LKAT system, disconnect head and signal output cables from the meter unit, then connect the 3 cables from the LKAT MUT.  Power up LKAT system.

- Using the adjustments on the LKAT MUT, set the corresponding ‘head’ signal for the desired current level ( typically just before the alarm trip point )

- Using MUT adjustments, increase or decrease signal to check alarm trip points via alarm LED on LKAT PCB as well as the relay trip LED on LKAT MUT.

( Similar method used to modify alarm trip points if needed )

# Ordering Information

046985	Tester, w/Adapters compatible with LKAT+ & 2

(contact factory for LKAT2 oem/sigmas with s/n 1403… to 1409…)